Western blotting

- Polyacrylamide gel   (modified from Molecular Cloning Volume 3)

|  | 5% | 6% | 8% | 10% | 12% | 15% |
| --- | --- | --- | --- | --- | --- | --- |
| Separation (KD) | Stacking | 57 – 212 | 36 - 94 | 20 - 80 | 12 - 60 | 10 - 43 |
| ddH2O | 5800 ul | 9240 ul | 8440 ul | 7640 ul | 6840 ul | 5640 ul |
| 40% acrylamide | 1000 ul | 2400 ul | 3200 ul | 4000 ul | 4800 ul | 6000 ul |
| 1.5 M Tris (pH8.8) |  | 4000 ul | 4000 ul | 4000 ul | 4000 ul | 4000 ul |
| 1.0 M Tris (pH6.8) | 1000 ul |  |  |  |  |  |
| 10% SDS | 80 ul | 160 ul | 160 ul | 160 ul | 160 ul | 160 ul |
| 10% APS | 80 ul | 160 ul | 160 ul | 160 ul | 160 ul | 160 ul |
| TEMED | 8 ul | 8 ul | 8 ul | 8 ul | 8 ul | 8 ul |
| Total | 8 ml | 16 ml | 16 ml | 16 ml | 16 ml | 16 ml |

- Preparation of gel (Model:Mini-PROTEAN Tetra Cell, Bio Rad)

- Assemble the glass plates and spacers (1.5 mm or 0.75 mm thick).

2.Pour the running gel(~10 ml).

3.Seal with 1 ml water-saturated 1-butanol.

4.When running gel has set, pour off the butanol and rinse with deionized water.

5.Pour the stacking gel (~5 ml) and insert the comb immediately.

6.When the stacking gel has set, place in gel rig and immerse in buffer.

7.Prior to running the gel, flush the wells out thoroughly with running buffer.

- Running the gel

- Add SDS-page sample buffer (5X) into the protein sample, and heat the mixture at 95℃ for 5 minutes.

- Let the mixture cool at room temperature for 10 minutes, and then spinning the samples.

- load the sample into the wells, and be sure to use markers.

- Run with constant current (Model:Mini-PROTEAN Tetra Cell, Bio Rad, voltage set at 110 V；Model:SE 600 Ruby, GE Healthcare, set at 50 mA/gel )Usual running time is about 2 hr.

- Preparation of membrane

- Cut a piece of PVDF membrane.

- Wet for about 5 min in methanol on a rocker at room temp. Remove methanol and add 1x transfer buffer until ready to use.

- Membrane transfer

- Assemble "sandwich" for Amersham 's Transblot (Model:TE 22).

- Pre-wet the sponges, 3M papers (slightly bigger than gel) in 1x transfer buffer. 
      Sponge - filter paper - gel - membrane - filter paper - sponge

- Transfer 100 V for 1.5-2 hr at 4°C.

- When finished, immerse membrane in Blocking buffer at room temp for 1 hr.

- Antibodies and detection

- Add primary antibody diluted in blocking buffer and incubated at 4°C O/N

- Wash membrane three times with PBST, 10 min each.

- Add secondary antibody diluted in blocking buffer at room temp for 1 hr.

- Wash membrane three times with PBST, 10 min each.

- Detect with Amersham ECL kit or ECL plus kit.

- Buffers for Westerns

- 5x sample buffer: (250 mM Tris-HCl, pH6.8, 400 mM DTT, 10%SDS, 0.025% bromophenol blue, 50% glycerol)

2.5 ml 1 M Tris-HCl, pH6.8

0.6175 g DTT

1 g SDS

0.0025g Bromophenol blue

5 ml glycerol

ddH2O  to 10 ml

- Resolving gel buffer: 100 ml

0.8 g SDS (add last)

18.2 g Trizma base (= 1.5 M)

Adjust pH to 8.8 with concentrated HCl

- Stacking gel buffer: 100 ml

0.4 g SDS (add last)

12.1 g Trizma base (= 1 M)

Adjust pH to 6.8 with concentrated HCl

- 10x Running buffer: 1 L

30.3 g Trizma base (= 0.25 M)

144 g Glycine (= 1.92 M)

10 g SDS (= 1%)--add last

Do not adjust the pH!!

- 10x transfer buffer: 1 L

58.1 g Trizma base (= 0.48 M)

29.3 g Glycine (= 0.39 M)

3.75 g SDS (=0.0375%)

To make 0.5 L of 1x transfer buffer:

100 ml Methanol

50 ml 10x transfer buffer

ddH2O to 500 ml

- 0.1% PBST Wash buffer 1 L

To make 1 L of 1x wash buffer:

100 ml 10x PBS

1 ml Tween 20

ddH2O  to 1000 ml

- Blocking buffer

5% non-fat milk

Make up in 0.1% PBST